## data summary
id1 | compound | n | amp2 mean [mV] | amp2 sd | true_d2H_pred mean [permil] | true_d2H_pred sd | true_d2H_pred_se mean [permil] | d2H rounded to integer [permil]
Bio 1.0 Mid | CH4 | 4 | 1371.4473 | 311.4895 | -240.1179 | 2.1877 | 6.6041 | -240
H-ISO1 | CH4 | 6 | 1401.8029 | 227.0071 | -162.5229 | 1.8294 | 6.0571 | -163
OM17-14C-BS2 | CH4 | 2 | 10259.2752 | 9560.5527 | -271.2984 | 0.82329 | 7.1867 | -271
OM17-4-DC | CH4 | 1 | 2499.0279 |  | -225.0899 |  | 6.3884 | -225
OM17-71A-BS1 | CH4 | 1 | 4526.8114 |  | -313.6977 |  | 8.2106 | -314
OM17-71A-BS2 | CH4 | 1 | 5139.8484 |  | -311.1817 |  | 8.1439 | -311
OM18-CM2A-BS1A | CH4 | 1 | 3224.9253 |  | -206.0512 |  | 6.1864 | -206
OM18-NSHQ14-BS1A | CH4 | 1 | 2997.6625 |  | -314.2516 |  | 8.2253 | -314
OM18-NSHQ4-BS1A | CH4 | 1 | 2438.2676 |  | -229.4136 |  | 6.4455 | -229
OM18-WAB71-BS1A | CH4 | 1 | 4897.8722 |  | -306.6163 |  | 8.0245 | -307
T-ISO1 | CH4 | 11 | 1375.1461 | 220.645 | -153.3992 | 4.8316 | 6.0933 | -153

## all data
id1 | is_std_peak | file_id | file_root | file_path | file_subpath | file_datetime | Row | Peak Center | Check Ref. Dilution | H3 Stability | H3 Factor | Conditioning | Seed Oxidation | GC Method | AS Method | inject_type | Analysis | Comment | Method | measurement_info | MS_integration_time.s | inject_volume [uL] | folder | compound | peak_info | ref_nr | peak_nr | is_ref | rt_start [s] | rt | rt_end [s] | amp2 [mV] | amp3 [mV] | bgrd2_start [mV] | bgrd3_start [mV] | bgrd2_end [mV] | bgrd3_end [mV] | area2 [mVs] | area3 [mVs] | r3/2 | rd3/2 [permil] | d3/2 [permil] | d2H [permil] | at2H [%] | n_overlapping | n_matches | is_identified | is_missing | is_ambiguous | true_d2H [permil] | in_calib | resid [permil] | true_d2H_pred [permil] | true_d2H_pred_se [permil] | in_range | calib | calib_ok | calib_points
T-ISO1 | TRUE | 1341__T-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1341__T-ISO1_1000.dxf |  | 2018-06-26 12:27 | 13 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1341 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547) |  | 5 | 0 | 531.487 | 546.953 | 612.579 | 1709.5582 | 522.0783 | 22.7075 | 138.4533 | 22.7075 | 138.4533 | 25079.3159 | 7198.7561 | 0.28704 | 276.7793 | 276.7793 | 276.7793 | 0.019882 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 13.9959 | -147.4932 | 6.1237 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1342__T-ISO1_0900.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1342__T-ISO1_0900.dxf |  | 2018-06-26 13:10 | 14 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1342 | 0900 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61892], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 900 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.323 | 547.162 | 606.518 | 1381.8423 | 416.6624 | 22.8027 | 139.2746 | 22.8027 | 139.2746 | 18239.5122 | 5184.2162 | 0.28423 | 264.5506 | 264.5506 | 264.5506 | 0.019691 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 1.7672 | -155.7996 | 6.0797 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1343__T-ISO1_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1343__T-ISO1_0800.dxf |  | 2018-06-26 13:23 | 15 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1343 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.532 | 547.162 | 601.92 | 1127.3026 | 337.1374 | 22.5278 | 139.2923 | 22.5278 | 139.2923 | 14187.4026 | 4042.7594 | 0.28495 | 268.5666 | 268.5666 | 268.5666 | 0.019754 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 5.7832 | -153.0717 | 6.0922 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1344__T-ISO1_0700.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1344__T-ISO1_0700.dxf |  | 2018-06-26 13:40 | 16 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1344 | 0700 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 700 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.4) |  | 5 | 0 | 532.532 | 547.371 | 593.56 | 1075.6413 | 322.2942 | 22.49 | 139.2679 | 22.49 | 139.2679 | 14038.0232 | 4025.1561 | 0.28673 | 272.9598 | 272.9598 | 272.9598 | 0.019822 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 10.1764 | -150.0876 | 6.1081 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1353__T-ISO1_0950.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1353__T-ISO1_0950.dxf |  | 2018-06-26 17:20 | 25 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1353 | 0950 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61892], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 950 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.114 | 547.162 | 605.055 | 1430.2971 | 431.1114 | 22.6657 | 137.8037 | 22.6657 | 137.8037 | 18516.6394 | 5286.8204 | 0.28552 | 271.604 | 271.604 | 271.604 | 0.019801 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 8.8205 | -151.0086 | 6.103 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1354__T-ISO1_0850.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1354__T-ISO1_0850.dxf |  | 2018-06-26 17:35 | 26 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1354 | 0850 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 850 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.323 | 547.162 | 604.01 | 1246.4567 | 372.0123 | 22.3226 | 138.3729 | 22.3226 | 138.3729 | 15766.6293 | 4491.3659 | 0.28487 | 267.5943 | 267.5943 | 267.5943 | 0.019739 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 4.8109 | -153.7322 | 6.089 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1355__T-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1355__T-ISO1_1000.dxf |  | 2018-06-26 17:48 | 27 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1355 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547) |  | 5 | 0 | 531.487 | 546.953 | 570.988 | 1570.0679 | 472.526 | 22.2549 | 137.8646 | 22.2549 | 137.8646 | 20323.0855 | 5745.6456 | 0.28272 | 255.7315 | 255.7315 | 255.7315 | 0.019554 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | -7.052 | -161.7901 | 6.0587 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1357__T-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1357__T-ISO1_1000.dxf |  | 2018-06-26 18:47 | 29 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1357 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.5) |  | 5 | 0 | 531.696 | 546.535 | 608.399 | 1557.8017 | 469.1694 | 22.2723 | 138.0051 | 22.2723 | 138.0051 | 20409.3037 | 5761.8269 | 0.28231 | 257.8596 | 257.8596 | 257.8596 | 0.019587 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | -4.9239 | -160.3446 | 6.0629 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1362__T-ISO1_0750.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1362__T-ISO1_0750.dxf |  | 2018-06-26 20:28 | 34 | 1 | 0 | 0 | 7.93812064911777 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | lin.std | 1362 | 0750 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 750 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.741 | 547.162 | 595.65 | 1053.0883 | 314.4591 | 22.3042 | 138.3528 | 22.3042 | 138.3528 | 12751.6503 | 3668.5947 | 0.2877 | 279.0821 | 279.0821 | 279.0821 | 0.019918 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 16.2986 | -145.929 | 6.1339 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1381__H-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1381__H-ISO1_1000.dxf |  | 2018-06-27 13:33 | 14 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1381 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61902], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.7) |  | 5 | 0 | 530.86 | 545.699 | 601.92 | 1583.6804 | 475.3944 | 22.1322 | 137.4991 | 22.1322 | 137.4991 | 20485.6813 | 5750.6344 | 0.28071 | 249.5431 | 249.5431 | 249.5431 | 0.019458 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -14.7126 | -165.9937 | 6.0494 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1382__H-ISO1_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1382__H-ISO1_0800.dxf |  | 2018-06-27 13:47 | 15 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1382 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.5) |  | 5 | 0 | 531.905 | 546.535 | 595.441 | 1200.2992 | 357.1142 | 22.0272 | 138.5112 | 22.0272 | 138.5112 | 14937.3935 | 4224.2104 | 0.28279 | 257.215 | 257.215 | 257.215 | 0.019577 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -7.0406 | -160.7824 | 6.0616 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1385__H-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1385__H-ISO1_1000.dxf |  | 2018-06-27 14:30 | 18 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1385 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.5) |  | 5 | 0 | 530.651 | 545.49 | 606.1 | 1594.4165 | 478.7721 | 22.0446 | 137.6888 | 22.0446 | 137.6888 | 20599.0896 | 5808.5302 | 0.28198 | 254.9332 | 254.9332 | 254.9332 | 0.019542 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -9.3225 | -162.3324 | 6.0572 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1386__H-ISO1_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1386__H-ISO1_0800.dxf |  | 2018-06-27 14:45 | 19 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1386 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61902], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.3) |  | 5 | 0 | 531.487 | 546.326 | 612.161 | 1207.1541 | 358.111 | 22.0008 | 138.2482 | 22.0008 | 138.2482 | 15231.2488 | 4304.1681 | 0.28259 | 256.4444 | 256.4444 | 256.4444 | 0.019565 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -7.8113 | -161.3059 | 6.06 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1389__H-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1389__H-ISO1_1000.dxf |  | 2018-06-27 15:35 | 22 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1389 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61892], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.7) |  | 5 | 0 | 530.86 | 545.699 | 606.936 | 1645.8238 | 494.9219 | 22.1474 | 138.1806 | 22.1474 | 138.1806 | 21514.5101 | 6061.1438 | 0.28172 | 254.58 | 254.58 | 254.58 | 0.019536 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -9.6756 | -162.5723 | 6.0566 | in range | delta_only | TRUE | 21
H-ISO1 | TRUE | 1390__H-ISO1_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1390__H-ISO1_0800.dxf |  | 2018-06-27 15:51 | 23 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1390 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61902], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.3) |  | 5 | 0 | 531.696 | 546.326 | 597.74 | 1179.4435 | 349.8154 | 22.0511 | 138.2074 | 22.0511 | 138.2074 | 14598.7425 | 4118.5888 | 0.28212 | 255.2002 | 255.2002 | 255.2002 | 0.019546 | 1 | 1 | TRUE | FALSE | FALSE | -156 | TRUE | -9.0554 | -162.151 | 6.0577 | in range | delta_only | TRUE | 21
Bio 1.0 Mid | TRUE | 1393__Bio 1_0 Mid_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1393__Bio 1_0 Mid_1000.dxf |  | 2018-06-27 17:20 | 26 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1393 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61892], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.9) |  | 5 | 0 | 530.651 | 545.908 | 612.788 | 1718.2619 | 469.5643 | 22.4516 | 137.6086 | 22.4516 | 137.6086 | 23825.4904 | 6098.4121 | 0.25596 | 141.4115 | 141.4115 | 141.4115 | 0.017774 | 1 | 1 | TRUE | FALSE | FALSE | -240 | TRUE | 0.81971 | -239.4432 | 6.5931 | in range | delta_only | TRUE | 21
Bio 1.0 Mid | TRUE | 1394__Bio 1_0 Mid_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1394__Bio 1_0 Mid_0800.dxf |  | 2018-06-27 22:55 | 27 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1394 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.5) |  | 5 | 0 | 532.323 | 546.535 | 602.338 | 1093.865 | 293.0495 | 23.2718 | 138.7401 | 23.2718 | 138.7401 | 13069.6533 | 3352.1095 | 0.25648 | 140.8376 | 140.8376 | 140.8376 | 0.017765 | 1 | 1 | TRUE | FALSE | FALSE | -240 | TRUE | 0.24578 | -239.8331 | 6.5992 | in range | delta_only | TRUE | 21
Bio 1.0 Mid | TRUE | 1397__Bio 1_0 Mid_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1397__Bio 1_0 Mid_1000.dxf |  | 2018-06-27 23:41 | 30 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1397 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61883], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.1) |  | 5 | 0 | 531.069 | 546.117 | 602.129 | 1550.5463 | 422.7854 | 22.709 | 137.8469 | 22.709 | 137.8469 | 19907.5067 | 5083.1101 | 0.25534 | 135.9067 | 135.9067 | 135.9067 | 0.017689 | 1 | 1 | TRUE | FALSE | FALSE | -240 | TRUE | -4.6851 | -243.1824 | 6.6533 | in range | delta_only | TRUE | 21
Bio 1.0 Mid | TRUE | 1398__Bio 1_0 Mid_0800.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1398__Bio 1_0 Mid_0800.dxf |  | 2018-06-28 00:01 | 31 | 1 | 0 | 0 | 8.03765609891788 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | test | 1398 | 0800 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61873], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 800 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.5) |  | 5 | 0 | 532.114 | 546.535 | 607.563 | 1123.1159 | 302.1583 | 22.6577 | 138.4868 | 22.6577 | 138.4868 | 13662.1789 | 3514.9005 | 0.25727 | 143.5172 | 143.5172 | 143.5172 | 0.017807 | 1 | 1 | TRUE | FALSE | FALSE | -240 | TRUE | 2.9254 | -238.0129 | 6.5708 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1470__T-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1470__T-ISO1_1000.dxf |  | 2018-07-02 11:07 | 4 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | drift.std | 1470 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61881], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.9) |  | 5 | 0 | 531.069 | 545.908 | 605.473 | 1532.0468 | 464.902 | 25.0459 | 139.2094 | 25.0459 | 139.2094 | 19816.6743 | 5635.0413 | 0.28436 | 266.3268 | 266.3268 | 266.3268 | 0.019719 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 3.5433 | -154.5931 | 6.085 | in range | delta_only | TRUE | 21
T-ISO1 | TRUE | 1471__T-ISO1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1471__T-ISO1_1000.dxf |  | 2018-07-02 11:21 | 5 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | drift.std | 1471 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61891], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.3) |  | 5 | 0 | 531.278 | 546.326 | 600.248 | 1442.5041 | 435.3552 | 24.1265 | 139.6684 | 24.1265 | 139.6684 | 18296.8781 | 5215.8595 | 0.28507 | 267.8755 | 267.8755 | 267.8755 | 0.019743 | 1 | 1 | TRUE | FALSE | FALSE | -157 | TRUE | 5.092 | -153.5412 | 6.09 | in range | delta_only | TRUE | 21
OM17-14C-BS2 | FALSE | 1472__OM17-14C-BS2_0600.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1472__OM17-14C-BS2_0600.dxf |  | 2018-07-02 12:01 | 6 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1472 | 0600 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61881], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 600 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (544.4) |  | 5 | 0 | 528.979 | 544.445 | 641.212 | 17019.6069 | 6651.2594 | 24.1824 | 137.9126 | 24.1824 | 137.9126 | 220112.1504 | 54148.4666 | 0.246 | 95.3717 | 95.3717 | 95.3717 | 0.017058 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -270.7163 | 7.1743 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM17-71A-BS1 | FALSE | 1475__OM17-71A-BS1_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1475__OM17-71A-BS1_1000.dxf |  | 2018-07-02 13:16 | 8 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1475 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61891], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.7) |  | 5 | 0 | 529.606 | 545.699 | 623.865 | 4526.8114 | 1244.2794 | 23.9916 | 138.4186 | 23.9916 | 138.4186 | 59869.1316 | 13886.4983 | 0.23195 | 32.095 | 32.095 | 32.095 | 0.016072 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -313.6977 | 8.2106 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM17-71A-BS2 | FALSE | 1476__OM17-71A-BS2_0770.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1476__OM17-71A-BS2_0770.dxf |  | 2018-07-02 13:33 | 9 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1476 | 0770 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61891], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 770 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (546.5) |  | 5 | 0 | 530.442 | 546.535 | 627.627 | 5139.8484 | 1436.6176 | 23.5228 | 137.6426 | 23.5228 | 137.6426 | 66387.9737 | 15440.0072 | 0.23257 | 35.799 | 35.799 | 35.799 | 0.01613 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -311.1817 | 8.1439 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM17-4-DC | FALSE | 1477__OM17-4-DC_0200.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1477__OM17-4-DC_0200.dxf |  | 2018-07-02 13:47 | 10 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1477 | 0200 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61891], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 200 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 531.696 | 547.162 | 611.116 | 2499.0279 | 707.4695 | 23.2837 | 138.6475 | 23.2837 | 138.6475 | 33107.778 | 8660.7922 | 0.26159 | 162.5424 | 162.5424 | 162.5424 | 0.018103 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -225.0899 | 6.3884 | >'amp2' range | delta_only | TRUE | 21
OM17-14C-BS2 | FALSE | 1478__OM17-14C-BS2_0170.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1478__OM17-14C-BS2_0170.dxf |  | 2018-07-02 14:07 | 11 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1478 | 0170 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61881], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 170 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547) |  | 5 | 0 | 531.487 | 546.953 | 625.537 | 3498.9436 | 961.7603 | 23.3212 | 137.7861 | 23.3212 | 137.7861 | 45942.5426 | 11294.356 | 0.24584 | 93.6576 | 93.6576 | 93.6576 | 0.017031 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -271.8806 | 7.1991 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM18-WAB71-BS1A | FALSE | 1481__OM18-WAB71-BS1A_1000.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1481__OM18-WAB71-BS1A_1000.dxf |  | 2018-07-02 15:30 | 14 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1481 | 1000 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61872], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 1000 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (545.5) |  | 5 | 0 | 528.979 | 545.49 | 638.077 | 4897.8722 | 1356.8786 | 23.5667 | 138.2287 | 23.5667 | 138.2287 | 73958.6985 | 17330.225 | 0.23432 | 42.52 | 42.52 | 42.52 | 0.016235 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -306.6163 | 8.0245 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM18-NSHQ4-BS1A | FALSE | 1483__OM18-NSHQ4-BS1A_0100.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1483__OM18-NSHQ4-BS1A_0100.dxf |  | 2018-07-02 16:13 | 16 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1483 | 0100 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61891], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 100 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.114 | 547.162 | 617.386 | 2438.2676 | 690.1666 | 23.192 | 138.278 | 23.192 | 138.278 | 30260.6217 | 7858.1331 | 0.25968 | 156.177 | 156.177 | 156.177 | 0.018004 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -229.4136 | 6.4455 | >'amp2' range | delta_only | TRUE | 21
OM18-NSHQ14-BS1A | FALSE | 1484__OM18-NSHQ14-BS1A_0180.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1484__OM18-NSHQ14-BS1A_0180.dxf |  | 2018-07-02 16:29 | 17 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1484 | 0180 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61881], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 180 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.4) |  | 5 | 0 | 532.114 | 547.371 | 618.013 | 2997.6625 | 775.9879 | 23.0505 | 137.9862 | 23.0505 | 137.9862 | 36737.4156 | 8510.0932 | 0.23165 | 31.2794 | 31.2794 | 31.2794 | 0.01606 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -314.2516 | 8.2253 | >'amp2' range, <'true_d2H_pred' range | delta_only | TRUE | 21
OM18-CM2A-BS1A | FALSE | 1485__OM18-CM2A-BS1A_0170.dxf | . | data/OM17-18_CH4_H_data/OM17-18_CH4_H_dxf_files/1485__OM18-CM2A-BS1A_0170.dxf |  | 2018-07-02 16:48 | 18 | 1 | 0 | 0 | 8.01539890033766 | 0 | 0 | OGL Testing\DBN\180624 GC CP 30m CH C1 gases SSL splitless 660s 27C isothermal injector 70C 0.6 ml min.tr2.meth | Disabled | sample | 1485 | 0170 | OGL Testing\DBN\180624 MS CP 30m H CH4 manual 660s acquisition BF 190-465s enter start.met | Set Sample Dilution to 0 [%], Sample dilution: 001 0, Set Reference Dilution to 84 [%], Activate HTC Mode..., Peak Center found at [61881], Sample Dilution found at 0 [], Set Sample Dilution to 0 [%], Sample dilution: 001 0 | 0.2 | 170 | OM17-18_CH4_H_dxf_files | CH4 | CH4 (547.2) |  | 5 | 0 | 532.114 | 547.162 | 615.923 | 3224.9253 | 968.1389 | 23.0145 | 138.076 | 23.0145 | 138.076 | 38546.6787 | 10333.9533 | 0.26809 | 190.5709 | 190.5709 | 190.5709 | 0.01854 | 1 | 1 | TRUE | FALSE | FALSE |  | FALSE |  | -206.0512 | 6.1864 | >'amp2' range | delta_only | TRUE | 21

## calib coefs
calib | calib_ok | calib_points | term | estimate | std.error | statistic | p.value | signif
delta_only | TRUE | 21 | (Intercept) | 493.917 | 10.1767 | 48.5341 | 2.1805e-21 | *** (p < 0.001)
delta_only | TRUE | 21 | true_d2H | 1.4722 | 0.057953 | 25.403 | 3.9643e-16 | *** (p < 0.001)

## calib summary
calib | calib_ok | calib_points | r.squared | adj.r.squared | sigma | statistic | p.value | df | logLik | AIC | BIC | deviance | df.residual
delta_only | TRUE | 21 | 0.9714 | 0.96989 | 8.6932 | 645.313 | 3.9643e-16 | 2 | -74.1603 | 154.3206 | 157.4541 | 1435.874 | 19

## calib range
calib | calib_ok | calib_points | term | units | min | max
delta_only | TRUE | 21 | amp2 | mV | 1053.0883 | 1718.2619
delta_only | TRUE | 21 | true_d2H_pred | permil | -243.1824 | -145.929